## Slide 1
Smart Infusion Pump
Software Simulator
Technical Design & Requirements Report
v4.0 Extended Edition
Flutter (iOS/Android)
Supabase + PostgreSQL
May 2026
Doctor
Nurse
Admin

## Slide 2
OVERVIEW
Presentation Agenda
01
Requirements
Functional · Non-Functional · Safety · Security · Validation
02
Three-Tier Architecture
Tools, Stack, Development Steps, DB Schema
03
Full Design Overview
Module structure, State machine, Routing, Audit log
04
UI/UX Overview
Design tokens, Screen inventory, Edge-case coverage
05
Validation & Traceability
Test layers, Test cases, Full traceability matrix

## Slide 3
SECTION 1 · REQUIREMENTS
Functional Requirements
Doctor Role
F-D-01
High
View/Add/Edit/Delete drugs in the library
F-D-02
Critical
Define 4 mandatory drug parameters (conc., rate, hard/soft limits)
F-D-03
High
Read-only access to own infusion logs
F-D-04
Med
Generate & export PDF Infusion Summary / Alarm Frequency reports
Nurse Role
F-N-01
Critical
Select drugs from active library only
F-N-04
Critical
Real-time hard/soft limit validation on all inputs
F-N-05
Critical
Start, Pause, Resume infusion controls
F-N-06
Critical
Emergency Stop — always visible, full-width, red
F-N-07
High
1 Hz real-time pump dashboard (rate, vol, ETA, battery)
F-N-08
Critical
Alarm panel: acknowledge non-critical / resolve critical
Admin: F-A-01 User Account CRUD (assign exactly one role) · F-A-02 System health metrics & audit log viewing

## Slide 4
SECTION 1 · REQUIREMENTS
Non-Functional Requirements
Performance
N-N-01 (Critical) Dashboard refresh ≥ 1 Hz, UI latency ≤ 100 ms
N-N-02 (Critical) FSM response to nurse input ≤ 500 ms
N-D-01 (High) PDF report gen ≤ 3 s for 1,000 entries
N-A-01 (Med) Log export ≤ 10 s for 100,000 entries
Usability
N-D-02 (High) ≤ 3 taps to any drug library function
N-D-03 (High) Real-time inline validation in Drug Editor
N-N-04 (High) Nurse UI: contrast ratio ≥ 4.5:1 (WCAG AA)
Dynamic type support, min touch target 44×44 pt

## Slide 5
SECTION 1 · REQUIREMENTS
Safety & Security Requirements
Safety Requirements
SF-D-01
Critical
Hard limits enforced by sim engine; violations blocked
SF-D-02
High
Soft limits: yellow banner + mandatory acknowledgement
SF-D-03
Med
Prevent deletion of drug in active infusion session
SF-N-01
Critical
Start button disabled if any hard limit violated
SF-N-02
Critical
Emergency Stop always visible, never covered by modals
SF-N-03
High
Battery warning at 20%; auto-stop + critical alarm at 5%
SF-A-01
Critical
Block deletion of last Admin account
Security Requirements
S-ALL-01
Critical
Auth required for every role screen (National ID + password)
S-D-01
Critical
RBAC: only Doctor edits drug library & generates reports
S-D-02
High
All Doctor actions logged in immutable audit log
S-N-01
Critical
Nurse barred from drug library, user mgmt, system config
S-A-01
Critical
Only Admin accesses user management
S-A-02
Critical
Admin password : bcrypt, cost factor ≥ 12, National ID is encrypted by sha256
S-A-03
High
Admin session timeout: 10 min (Doctor/Nurse: 15 min)

## Slide 6
SECTION 2 · ARCHITECTURE
Three-Tier Architecture Model
Tier 1
Presentation
Flutter 3.x + GoRouter 14.x
Role-isolated screen trees
Role badge in AppBar always visible
Session timeout banners
Tier 2
Business Logic
Riverpod 2.5 StateNotifier
Infusion FSM (10 states)
LimitValidator · BatterySimulator
AuditLogger dispatch
PDF Report Generator
RBAC guard logic
Tier 3
Data
Supabase (Auth + PostgreSQL + Realtime + Storage)
Drift SQLite — offline fallback
Row-Level Security on every table
AES-256 at rest · TLS in transit
Immutable audit_log (RLS block UPDATE/DELETE)

## Slide 7
SECTION 2 · ARCHITECTURE
Technology Stack
Layer
Technology
Purpose
UI Framework
Flutter 3.x
Cross-platform iOS & Android from single Dart codebase
State Mgmt
Riverpod 2.5 + riverpod_annotation
Reactive, compile-safe provider hierarchy
Routing
GoRouter 14.x
Declarative routing with RBAC redirect guard
Backend
Supabase
Auth, PostgreSQL, Realtime WebSocket, Storage
Offline DB
Drift (SQLite)
On-device fallback for simulation continuity
Models
freezed + json_serializable
Immutable, copyWith-able, JSON-serialisable domain models
PDF Export
pdf 3.11 + printing 5.13
Multi-page clinical reports from Doctor role
Security
flutter_secure_storage + local_auth
OS keychain token storage · Biometric auth

## Slide 8
SECTION 2 · ARCHITECTURE
Development Build Sequence (10 Steps)
01
Auth Shell
Login UI → Supabase Auth → role-routing after login
02
Role Router
GoRouter ShellRoutes · RBAC redirect · AccessDeniedScreen
03
Nurse UI Shell
PumpDashboard · Controls · EmergencyStop · InfusionNotifier
04
Nurse Logic
BatterySimulator · LimitValidator · KVOHandler · FSM transitions
05
Doctor UI Shell
DrugLibrary · DrugEditor → Supabase drug table
06
Doctor Logic
LimitValidator · AuditLogger · ReportGenerator (PDF)
07
Admin UI Shell
StatusDashboard · UserManagement · LogsViewer
08
Integration
Supabase Realtime alarm streaming · RLS policy testing
09
PDF Export
Infusion Summary & Alarm Frequency PDFs from filtered data
10
Polish
Session timeout · Offline Drift sync · WCAG audit · Edge cases

## Slide 9
SECTION 2 · ARCHITECTURE
Database Schema — PostgreSQL (Supabase)
Table
PK
Description
users
user_id (UUID)
All accounts · soft-delete · national_id, password_hash, role, last_login
drug
drug_id
Drug library managed by Doctors · concentration, rates, hard/soft limits
infusion_session
session_id
Full parameter snapshot per session · status enum · volume accounting
alarms
alarm_id
Reference table of all alarm types with severity classifications
alarm (events)
event_id
Instance log per session · ack/resolve tracking with user + timestamp
audit_log
log_id
IMMUTABLE append-only · old_value/new_value · PostgreSQL rules block UPDATE & DELETE
Row-Level Security (RLS) policies enforced on every table — second line of defence independent of Flutter UI RBAC guards

## Slide 10
SECTION 3 · DESIGN OVERVIEW
Feature-First Module Organisation
lib/core/
router/ — GoRouter + RBAC guard
theme/ — Role ThemeData tokens
constants/ — timeouts, thresholds
error/ — sealed AppException
lib/features/auth/
LoginScreen · AuthRepository
Supabase auth datasource
AppUser model · RoleType enum
lib/features/doctor/
DrugRepository · ReportRepository
Drug / Report freezed models
Use cases: add / edit / delete / generate
DrugLibrary, DrugEditor, Logs, Reports
lib/features/nurse/
InfusionStateMachine (10 states)
BatterySimulator · LimitValidator · KVOHandler
All Nurse screens + AlarmPanel
InfusionState / AlarmType enums
lib/features/admin/
UserRepository · SystemHealthRepository
CRUD use cases · Admin screens
ManagedUser model
lib/shared/
RoleBadge · ConfirmationDialog
SessionTimeoutBanner
AccessDeniedScreen · SessionProvider

## Slide 11
SECTION 3 · DESIGN OVERVIEW
Simulation Engine & Infusion State Machine
IDLE
PROGRAMMING
RUNNING
PAUSED
ALARM
KVO
COMPLETE
EMRG STOP
BATT LOW
InfusionStateMachine
10-state FSM: idle → programming → running → paused → alarm → kvo → complete → emergencyStop → batteryLow → criticalBattery
Volume: rate/3600 mL/s per tick
Supabase sync throttled: every 60 s
BatterySimulator
Parallel drain loop: −0.1%/s during Running/KVO
@20% → low-battery alarm (non-blocking)
@5% → calls emergencyStop() + critical alarm
LimitValidator
Hard limit breach → input rejected, Start disabled
Soft limit breach → yellow warning, acknowledgement required
Called from both UI layer and use-case layer

## Slide 12
SECTION 3 · DESIGN OVERVIEW
Security Architecture & Audit Logging
Authentication
Supabase Auth · National ID + password · Biometric (local_auth) optional 2FA
Token Storage
flutter_secure_storage — OS Keychain (iOS) / Keystore (Android) — no plaintext
Transport
TLS (HTTPS) enforced — no plaintext connections to Supabase permitted
Data at Rest
AES-256 encryption on all stored logs and sensitive data via Supabase
Row-Level Security
Supabase RLS policies — authoritative DB-layer access control, independent of UI RBAC
Session Timeout
Admin: 10 min · Doctor / Nurse: 15 min · 60 s countdown warning banner
Key Audit Event Types
DRUG_ADDED / DRUG_EDITED / DRUG_ARCHIVED
Doctor
INFUSION_STARTED / PAUSED / STOPPED / EMERGENCY_STOP
Nurse
ALARM_ACKNOWLEDGED
Nurse
USER_CREATED / USER_DEACTIVATED / CONFIG_CHANGED
Admin

## Slide 13
SECTION 4 · UI/UX OVERVIEW
Design System & Token Specification
Doctor
#00685D
Nurse
#005EA4
Admin
#6200EA
Emrg.
#E53935
Start
#2E7D32
Warn
#F57F17
Typography Tokens
dataDisplay
32pt
Bold 700
Real-time numeric values (rate, volume, uptime %)
titleStyle
20pt
SemiBold 600
Screen titles, card headings
bodyStyle
16pt
Regular 400
Body text, descriptions, helper text
captionStyle
13pt
Medium 500
Input labels, secondary metadata
labelCaps
11pt
Bold 700 +tracked
ALL-CAPS nav/button labels
Spacing & Radius
spaceXs
4 dp
spaceSm
8 dp
spaceMd
16 dp
spaceLg
24 dp
radiusDefault
2 dp
radiusLg
4 dp
nurseRadiusXl
12 dp
nurseRadiusFull
9999 dp

## Slide 14
SECTION 4 · UI/UX OVERVIEW
Screen Inventory by Role
Doctor
Drug Library — searchable list, Add FAB, swipe edit/delete
Drug Editor — sectioned form (IDENTIFICATION, POTENCY, FLOW CONTROL), inline validation
Logs Viewer — filterable, read-only, tap for detail
Reports — type selector, date filter, preview pane, Export PDF
Nurse
Drug Selection — active drugs only, tap to proceed
Parameter Entry — Conc., Rate, VTBI, dosing calc panel, Emergency Stop always visible
Pump Dashboard — 2×2 data cards at 32pt bold, VTBI progress bar, 1 Hz refresh
Pump Controls — Start/Pause/Resume grid + Emergency Stop (2-sec hold + haptic)
Active Alarms — severity indicators, critical locks until resolved
Session Log Preview — read-only chronological events
Admin
Admin Dashboard — 2×2 stat cards, Active Alerts with coloured borders
User Management — role-coloured avatars, role pills, Add FAB, deactivate
Logs Viewer — date/drug/user filters, paginated table with event-type chips

## Slide 15
SECTION 4 · UI/UX OVERVIEW
Edge Case & Error State Coverage
Auth Role
Wrong credentials → red banner #FFEBEE above inputs
Deactivated account → 'Account has been deactivated'
No internet → 'No connection — cannot authenticate'
Biometric failed → fallback to password, no crash
Doctor Role
Empty drug library → empty state: icon + Add button
Delete drug in active session → blocking error modal
Hard Limit < Soft Limit → inline validation error
PDF export failed → snackbar with retry option
Nurse Role
Hard limit exceeded → Start disabled + blocking modal
Soft limit exceeded → yellow banner, 'I understand — proceed'
Critical alarm mid-infusion → pump auto-pauses, alarm panel overlays
Battery @20% → non-blocking banner; @5% → auto-stop + critical alarm
Emergency Stop accidental press → 2-sec hold + haptic progress fill
Admin Role
Delete last Admin → 'Create backup account first' error
Deactivate own account → warning dialog
Memory >80% → progress bar turns amber
Storage >90% → red left-border alert row

## Slide 16
SECTION 5 · VALIDATION & TESTING
Multi-Layer Testing Strategy
Unit Tests
LimitValidator · BatterySimulator · InfusionStateMachine · AuditLogger · PDF generator
Flutter test + Mockito for Supabase mocks
100% coverage target for safety-critical validators
Widget Tests
DrugEditor validation feedback, ParameterEntry limit warnings, AlarmPanel acknowledgement
flutter_test · pump.pumpAndSettle()
Golden tests for visual regression
Integration Tests
End-to-end role flows vs. Supabase test project with seeded data
flutter_driver / integration_test package
CI pipeline on GitHub Actions
RBAC / Security
Verify no cross-role navigation · RLS policy enforcement
Session timeout · last-Admin deletion block
Automated integration tests + manual RLS inspection
Performance
Dashboard 1 Hz refresh latency · FSM response time
Report generation speed · log export throughput
Dart DevTools timeline · manual stopwatch on target hardware
Accessibility
Contrast ratios ≥ 4.5:1 · touch targets ≥ 44×44 pt
VoiceOver / TalkBack screen reader compatibility
Flutter Accessibility Analyser + physical device testing

## Slide 17
SECTION 5 · VALIDATION & TESTING
Verification Test Cases
ID
Role
Type
Test Description
Expected Result
V-D-01
Doctor
Functional
Add drug with valid limits
Drug appears in library; selectable by Nurse role
V-D-02
Doctor
Validation
Set hard limit = −5 mg/mL
Inline validation error; data not saved
V-D-03
Doctor
Audit
Edit drug; open audit trail
Entry: user ID, timestamp, field, old val, new val
V-D-04
Doctor
Safety
Delete drug in active simulation
System blocks with explanatory error message
V-N-01
Nurse
Functional
Program valid infusion; start
Volume counter ↑; time remaining ↓ at 1 Hz
V-N-02
Nurse
Safety
Rate above hard limit
Start disabled; blocking error shown
V-N-03
Nurse
Safety
Rate above soft limit
Yellow warning banner; acknowledgement required
V-N-04
Nurse
Alarm
Trigger occlusion alarm mid-infusion
Auto-pause; critical alarm locked until resolved
V-N-05
Nurse
Emergency
Press Emergency Stop (2-sec hold)
Pump shuts down ≤ 500 ms; logged in audit trail
V-A-01
Admin
Functional
Create Nurse account; log in
Nurse-only screens accessible; Admin returns 403
V-A-02
Admin
Safety
Delete last Admin
Blocks: 'Only one Admin remaining. Create backup.'
V-A-03
Admin
Audit
5 Admin config changes; review log
All 5 entries with user ID, timestamp, changed vals

## Slide 18
SECTION 5 · VALIDATION & TESTING
Requirements Traceability Matrix
Req ID
Category
Requirement Summary
Verification Test(s)
Status
F-D-01
Functional
Doctor: view/add/edit/delete drugs
V-D-01
Defined
F-D-02
Functional
Doctor: define 4 mandatory drug params
V-D-01, V-D-02
Defined
F-N-04
Functional
Nurse: hard/soft limit validation
V-N-02, V-N-03
Defined
F-N-05
Functional
Nurse: Start/Pause/Resume
V-N-01
Defined
F-N-06
Functional
Nurse: Emergency Stop
V-N-05
Defined
F-N-07
Functional
Nurse: 1 Hz real-time dashboard
V-N-01, Perf Test
Defined
F-N-08
Functional
Nurse: alarm acknowledge/resolve
V-N-04
Defined
F-A-01
Functional
Admin: user account CRUD
V-A-01
Defined
N-N-01
Performance
Dashboard ≥ 1 Hz, latency ≤ 100 ms
Performance Test
Defined
N-N-02
Performance
FSM response ≤ 500 ms
V-N-05, Perf Test
Defined
S-ALL-01
Security
Auth required for all role screens
V-D-05, V-N-06, V-A-04
Defined
S-D-01
Security
RBAC: only Doctor accesses drug lib
V-D-05
Defined
S-A-02
Security
Admin passwords: bcrypt/Argon2 cost 12
Code Review + Sec Audit
Defined
SF-D-01
Safety
Hard limits enforced by sim engine
V-N-02
Defined
SF-D-02
Safety
Soft limit: yellow banner + ack
V-N-03
Defined
SF-N-02
Safety
Emergency Stop always visible
V-N-05, Access Test
Defined
SF-N-03
Safety
Battery @20% warn; @5% auto-stop
Unit Test (BatterySimulator)
Defined
SF-A-01
Safety
Last Admin deletion blocked
V-A-02
Defined

## Slide 19
Thank You
Smart Infusion Pump Software Simulator
Technical Design & Requirements Report v4.0
38
Requirements across 5 categories
10
FSM states in simulation engine
6
Test layers for full assurance
12
Verification test cases with traceability
Flutter · Riverpod · GoRouter · Supabase · PostgreSQL · Drift · AES-256